## Version 1.1 (Public)
Gamma Vacuum MPCq/QPC Controller Modbus Register Map
Document version: | 1.3
Document number: | 00.142.404
Date modified: | 2017-01-04
Firmware Version: | 1.16 or higher
Item |  | Quantity Allocated | Quantity Used | Number of Registers (16-bit) | Data Type | Supported Modbus Functions | Starting Item | Ending Item | Relative Hex Address | SPCE Controller Supported | Notes
HV On/Off |  | 8 | 4 | 0 | Boolean | 1, 5, 15 | 1 | 8 | 0x0000 | supported | 0 = turn off HV, 1 = turn on HV
(Reserved) |  | 8 | 4 | 0 |  | 1, 5, 15 | 9 | 16 | 0x0008 |  | Do not use.
(Reserved) |  | 1 | 1 | 0 |  | 1, 5, 15 | 17 | 17 | 0x0010 |  | Do not use.
(Reserved) |  | 8 | 4 | 0 |  | 1, 5, 15 | 18 | 25 | 0x0011 |  | Do not use.
Monitor Normal/Inverted |  | 8 | 4 | 0 | Boolean | 1, 5, 15 | 26 | 33 | 0x0019 |  | 0 = normal, 1 = inverted
(Reserved) |  | 1 | 1 | 0 |  | 1, 5, 15 | 34 | 34 | 0x0021 |  | Do not use.
Foldback On/Off |  | 8 | 4 | 0 | Boolean | 1, 5, 15 | 35 | 42 | 0x0022 | supported | 0 = foldback off, 1 = foldback on
System restart |  | 1 | 1 | 0 | Boolean | 1, 5, 15 | 43 | 43 | 0x002A | supported | Setting to 1 will restart the system. Reading always returns 0.
(Reserved) |  | 1 | 1 | 0 |  | 1, 5, 15 | 44 | 44 | 0x002B |  | Do not use.
Digital Input State |  | 8 | 4 | 0 | Boolean | 2 | 10001 | 10008 | 0x0000 |  | Readout only. 0 = default input state, not asserted. 1 = input state asserted.
Digital Output State (Relay) |  | 8 | 4 | 0 | Boolean | 2 | 10009 | 10016 | 0x0008 |  | Readout only. 0 = output state low, 1 = output state high.
Digital Output State (TTL) |  | 8 | 4 | 0 | Boolean | 2 | 10017 | 10024 | 0x0010 |  | Readout only . 0 = relay not activated. 1 = relay activated.
System Status |  | 512 | 256 | 0 | Boolean | 2 | 10025 | 10536 | 0x0018 | supported | Readout only. 64-bit field (See 'Display Error/Status Code' section in the user manual. Field bits correspond to status codes listed in the user manual. For example, if error 2 - vacuum loss is detected, bit 2 is asserted. ) Supported starting with firmware version 1.27.
(Reserved) |  | 8 | 4 | 8 |  | 4 | 30001 | 30064 | 0x0000 |  | Do not use.
Running Voltage |  | 8 | 4 | 1 | Integer | 4 | 30065 | 30072 | 0x0040 | supported | Supply output voltage in volts.
Running Current |  | 8 | 4 | 2 | IEEE754 | 4 | 30073 | 30088 | 0x0048 | supported | Supply output current in amps.
Running Pressure |  | 8 | 4 | 2 | IEEE754 | 4 | 30089 | 30104 | 0x0058 | supported | Pressure value in system pressure units.
(Reserved) |  | 8 | 4 | 1 |  | 4 | 30105 | 30112 | 0x0068 |  | Do not use.
Model Name |  | 1 | 1 | 8 | Integer | 4 | 30113 | 30120 | 0x0070 | supported | ASCII text of the controller model.
Firmware Version |  | 1 | 1 | 3 | Integer | 4 | 30121 | 30123 | 0x0078 | supported | Controller firmware version installed.
Ethernet MAC |  | 1 | 1 | 4 | Hex | 4 | 30124 | 30127 | 0x007B | supported | Ethernet MAC in hex digits.
(Reserved) |  | 1 | 1 | 1 |  | 4 | 30128 | 30128 | 0x007F |  | Do not use.
Modbus Status |  | 1 | 1 | 1 | Integer | 4 | 39999 | 39999 | 0x270E |  | 0=no error, 1=command pending, 2=data range error, 99=unknown error
Pump Size |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40001 | 40008 | 0x0000 | supported | Pump size in l/s.
Pump Pressure Factor |  | 8 | 4 | 2 | IEEE754 | 3, 6, 16 | 40009 | 40024 | 0x0008 |  | Default = 1.0. See user manual, pressure calculations section.
Auto HV Restart Control |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40025 | 40032 | 0x0018 |  | 0 = Off, 1 = Auto HV Restart On.
Pump Name |  | 8 | 4 | 8 | ASCII | 3, 6, 16 | 40033 | 40096 | 0x0020 | supported | ASCII Text of pump name. Maximum size 16 characters. Last character must be null character '\0'.
Voltage Target |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40097 | 40104 | 0x0060 | supported | Supply voltage target in volts.
Current Limit (mA) |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40105 | 40112 | 0x0068 | supported | Supply current limit in milliamps.
(Reserved) |  | 8 | 4 | 1 |  | 3, 6, 16 | 40113 | 40120 | 0x0070 |  | Do not use.
Foldback Pressure |  | 8 | 4 | 2 | IEEE754 | 3, 6, 16 | 40121 | 40136 | 0x0078 | supported | Foldback pressure in system pressure units.
Foldback Voltage |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40137 | 40144 | 0x0088 | supported | Foldback voltage in volts.
Analog Output Mode |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40145 | 40152 | 0x0090 |  | See user manual, analog output section.
Analog Output Supply |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40153 | 40160 | 0x0098 |  | See user manual, analog output section.
Analog Output Offset |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40161 | 40168 | 0x00A0 |  | See user manual, analog output section.
Digital Input Mode |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40169 | 40176 | 0x00A8 |  | 0=Off, 1=HV Enable, 2=HV Switch, 3=TSP Interlock1, 4=TSP Interlock2
Digital Input Supply |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40177 | 40184 | 0x00B0 |  | Supply (1-4)
Digital Output Mode (Relay) |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40185 | 40192 | 0x00B8 |  | 0=Off, 1=Pressure, 2=Error, 3=HV Indication, 4=Vacuum Loss
Digital Output Mode (TTL) |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40193 | 40200 | 0x00C0 |  | 0=Off, 1=Pressure, 2=Error, 3=HV Indication, 4=Vacuum Loss
Digital Output Supply (Relay) |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40201 | 40208 | 0x00C8 |  | Supply (1-4)
Digital Output Supply (TTL) |  | 8 | 4 | 1 | Integer | 3, 6, 16 | 40209 | 40216 | 0x00D0 |  | Supply (1-4)
Digital Output Pressure On (Relay) |  | 8 | 4 | 2 | IEEE754 | 3, 6, 16 | 40217 | 40232 | 0x00D8 |  | Pressure trip value in system pressure units.
Digital Output Pressure On (TTL) |  | 8 | 4 | 2 | IEEE754 | 3, 6, 16 | 40233 | 40248 | 0x00E8 |  | Pressure trip value in system pressure units.
Digital Output Pressure Off (Relay) |  | 8 | 4 | 2 | IEEE754 | 3, 6, 16 | 40249 | 40264 | 0x00F8 |  | Pressure trip value in system pressure units.
Digital Output Pressure Off (TTL) |  | 8 | 4 | 2 | IEEE754 | 3, 6, 16 | 40265 | 40280 | 0x0108 |  | Pressure trip value in system pressure units.
Pressure Units |  | 1 | 1 | 1 | Integer | 3, 6, 16 | 40281 | 40281 | 0x0118 |  | 0=Torr, 1 = mBar, 2=Pascal
Notes: | Data is big-endian. Data access and data storage are big-endian. Word-wise is big-endian. The most significant word of multi-word transactions are stored in lower memory location and delivered/receieved first in the serial stream. Byte-wise is big-endian. The most-significant byte of each word is stored in the lowest memory location, and is delivered/received first in the Modbus serial stream.
If the "Quantity Used" is larger than 1, then multiple registers are used to store the data item for separate system items. For example, the pump size for pump 1 is specified by register 40001 (address 0x0000), and the pump size for pump 2 is specified in register 40002 (address 0x0001), and so on